# FETCH subsystem
The FETCh subsystem reads the last measured values. It doesn’t trigger any measurements or clears any values.

### FETCh[:SCALar]:*[:DC]?
Syntax	FETCh[:SCALar]:CURRent[:DC]? / FETCh[:SCALar]:VOLTage[:DC]? / FETCh[:SCALar]:POWer[:DC]?
Description	These queries return the last measured current (A), voltage (V) or power (W) measurements. Measurements are averaged over the set number of line power cycles.
Return	Float
Usage example	Fetch the last measured current value: / FETCH:CURR? / 1.02
Errors
Related Commands	[SENSe:]NPLCycles

### FETCh[:SCALar]:*[:DC]:STATistics?
Syntax	FETCh[:SCALar]:CURRent[:DC]:STATistics? <stat> / FETCh[:SCALar]:VOLTage[:DC]:STATistics? <stat> / FETCh[:SCALar]:POWer[:DC]:STATistics? <stat>
Description	These queries return one or all of the statistical values of current (A), voltage (V) or power (W) measurements over the reported number amount of measurements (count). / Without parameters all stats of the measurement are retuned in the order min,avg,dev, max and count.
Parameters	Name	Type	Range	Default
	<stat>	Discrete	MIN|MAX|AVG|DEV|COUNT	all
Return	Float or comma separated floats. The count is an integer.
Usage example	Fetch the statistics of current: / FETCH:CURR:STAT? / 0.97,1.01,0.021,1.12,123
Errors
Related Commands	[SENSe:]NPLCycles / [SENSe:]CURRent:STATistics[:ON] / [SENSe:]CURRent:STATistics:CLEAR / [SENSe:]VOLTage:STATistics[:ON] / [SENSe:]VOLTage:STATistics:CLEAR / [SENSe:]POWer:STATistics[:ON] / [SENSe:]POWer:STATistics:CLEAR

### FETCh[:SCALar]:VOLTage:REVerse[:POLarity]?
Syntax	FETCh[:SCALar]:VOLTage:REVerse[:POLarity]?
Description	This query returns true if the last detected status of the reverse polarity detection circuit was positive (polarity reversed) or not (normal polarity).
Return	1 = reverse polarity, 0 = no reverse polarity
Usage example	FETCH:VOLT:REV? / 0
Errors
Related Commands

### FETCh:CAPacity?
Syntax	FETCh:CAPacity?
Description	This queries returns the last measured set of capacity values in the order Ah, Wh, time.
Return	Comma separated array
Usage example	FETCH:CAP? / 0.121, 0.232, 63
Errors
Related Commands	[SOURce]:CAPacity[:ON]