# SOURCE subsystem
The SOURce subsystem controls the input state, mode, levels and protections. These are are the primary functions and basic operation of the electronic load.
## [SOURce:]INPut subsystem
The INPUT subsystem controls the ON/OFF state, mode and general protection settings.
### [SOURce:]INPut[:STATe]
Syntax	[SOURce:]INPut[:STATe] {<bool>} / [SOURce:]INPut[:STATe]?
Description	This command enables or disables the input. The enabled state is ON (1); the disabled state is OFF (0). The state of a disabled output is a condition of highest impedance. / Self-test operation initiated by *TST? command will put the input into disable state. / Execution of the INP ON command on the channel which has one or more protection tripped (OCP, OVP, OPP, OTP or Reverse Polarity) will generate an error. / Use INPut:PROTection:CLEar command to clear all tripped protections.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	The query command returns 0 if the output is OFF, and 1 if the output is ON.
Usage example	Set and check output on: / INP ON / INP? / 1
Errors
Related Commands	*TST / INPut:PROTection:CLEar / [SOURce:]VOLTage:SENSe[:SOURce]

### [SOURce:]INPut:MODE
Syntax	[SOURce:]INPut:MODE {<mode>} / [SOURce:]INPut:MODE?
Description	This command programs the input mode. The mode can only be changed when the input state is off. / The modes are: / Constant Current (CC) / Constant Voltage (CV) / Constant Power (CP) / Constant Resistance (CR) / Volt Meter (DVM) / Short inputs (SHORT) / *RST resets the mode to CC.
Parameters	Name	Type	Range	Default
	<mode>	Discrete	CC|CV|CR|CP|DVM|SHORT	CC
Return	Query the input mode presently active.
Usage example	INP:MODE? / CC / INP:MODE CP / POWER 10 / INP 1 / FETCH:POW? / 10
Errors
Related Commands	INP / *RST / *TST

### [SOURce:]INPut:PROTection:CLEar
Syntax	[SOURce:]INPut:PROTection:CLEar
Description	This command clears the general protection status and the clears the individual states of OCP, OVP, OTP, OPP, Reverse Polarity and Sense Error error states.
Return	The command doesn’t return anything.
Usage example	INP:PROT:CLEar
Errors
Related Commands

### [SOURce:]INPut:PROTection:TRIPped
Syntax	[SOURce:]INPut:PROTection:TRIPped? / [SOURce:]INPut:PROTection:TRIPped:REVerse[:POLarity]?
Description	The TRIPped? query returns the general protection state. If any of OCP, OVP, OTP, OPP, Reverse polarity or Sense error are triggered during ON state the general protection is set and the load is turned off. / The TRIPped:REVerse query returns the state of the reverse polarity error state. This state is only triggered if reveres polarity is detected when the state is ON.
Return	This query returns 0 if a protection is not set, or 1 when it was not detected.
Usage example	INP:PROT:TRIP? / 1
Errors
Related Commands	[SOURce:]INPut:PROTection:CLEar

### [SOURce:]INPut:PROTection:WDOG
Syntax	[SOURce:]INPut:PROTection:WDOG {<bool>} / [SOURce:]INPut:PROTection:WDOG?
Description	This command enables or disables the SCPI activity watchdog. The SCPI activity watchdog implements a safeguard against a fault situation when SCPI connection with a host computer is lost and prevents the load from staying in the ON state in this situation. / If the SCPI watchdog triggers the protection state is set and the load is turned of. / The SCPI watchdog has two distinct ways to determine if the SCPI connection is still active: / - Based on any SCPI activity received / - Based on explicit PET commands
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	No return
Usage example	Set the SCPI watchdog type to PET, timeout to 10 seconds and activate the watchdog and pet the watchdog for the first time. / INP:WDOG:TYP PET / INP:WDOG:DEL 10 / INP:WDOG 1 / INP:WDOG:PET
Related Commands	[SOURce:]INPut:PROTection:WDOG:CLEar / [SOURce:]INPut:PROTection:WDOG:DELay / [SOURce:]INPut:PROTection:WDOG:PET / [SOURce:]INPut:PROTection:WDOG:TYPe

### [SOURce:]INPut:PROTection:WDOG:CLEar/TRIPped?
Syntax	[SOURce:]INPut:PROTection:WDOG:CLEar / [SOURce:]INPut:PROTection:WDOG:TRIPped?
Description	This command clears the latched protection status that disables the output when a SCPI watchdog timeout is reached. Note the watchdog is NOT turned of and will trigger again if activity or PET commands are not detected within the set delay. / The query returns whether the SCPI watchdog condition is set.
Usage example	The following command clears the SCPI watchdog status: / INP:WDOG:CLE
Related Commands	[SOURce:]INPut:PROTection:WDOG / [SOURce:]INPut:PROTection:TYPe / [SOURce:]INPut:PROTection:DELay

### [SOURce:]INPut:PROTection:WDOG:DELay
Syntax	[SOURce:]INPut:PROTection:WDOG:DELay {<delay>} / [SOURce:]INPut:PROTection:WDOG:DELay?
Description	This command set the SCPI watchdog timeout delay in seconds. / The query returns the current set delay in seconds.
Parameters	Name	Type	Range	Default
	<delay>	Integer	0-3600	10
Return	The query command returns the number of seconds the delay is set to, independent of the status of the watchdog.
Usage example	INP:WDOG:DEL 5 / INP:WDOG:DEL? / 5
Related Commands	INP:WDOG

### [SOURce:]INPut:PROTection:WDOG:PET
Syntax	[SOURce:]INPut:PROTection:WDOG:PET
Description	This command clears the SCPI watchdog inactivity timer.
Return	None
Usage example	INP:PROT:WDOG:PET

### [SOURce:]INPut:PROTection:WDOG:TYPe
Syntax	[SOURce:]INPut:PROTection:WDOG:TYPe {<wdogType>} / [SOURce:]INPut:PROTection:WDOG:TYPe?
Description	This command set the SCPI watchdog type to one of: / Activity: The SCPI watchdog timer is reset on reception of any SCPI command line / PET: The SCPI watchdog timer is only reset on the reception of the INPut:WDOG:PET command.
Parameters	Name	Type	Range	Default
	<wdogType>	Discrete	ACTivity|PET	ACT
Return	The type of SCPI watchdog type. ACT for Activity or PET for PET.
Usage example	Set the SCPI watchdog type to PET, set the delay to 5s, activate the SCPI watchdog: / INP:WDOG:TYP PET / INP:WDOG:DEL 5 / INP:WDOG 1 / Send the PET command at least once every 5 seconds: / INP:WDOG:PET / INP:WDOG:PET / Disable the SCPI watchdog when the critical routine is finished or when the remote program ends: / INP:WDOG 1
Errors
Related Commands	[SOURce:]INPut:WDOG

## CURRENT subsystem
### [SOURce:]CURRent[:LEVel][:IMMediate][:AMPLitude]
Syntax	[SOURce:]CURRent[:LEVel][:IMMediate][:AMPLitude] {<current>} / [SOURce:]CURRent[:LEVel][:IMMediate][:AMPLitude]?
Description	This command sets the desired current level in CC mode. The allowed range depends on the current range selected. / The query returns the set current level.
Parameters	Name	Type	Range	Default
	<current>	NR2	0.0-10.0	0.1
	<low range>	NR2	0.0-1.0	0.1
Return	Returns the set current level as float in Amperes.
Usage example	Set the mode to CC, set the current level to 1.5A and turn the load on: / INP:MODE CC / CURR 1.5 / INP 1
Errors
Related Commands	[SOURce:]INPut:MODE / [SOURce:]CURRENT:RANGe

### [SOURce:]CURRent:PROTection[:LEVel]
Syntax	[SOURce:]CURRent:PROTection[:LEVel] {<current>} / [SOURce:]CURRent:PROTection[:LEVel]?
Description	This command sets the desired over current protection (OCP) level. The allowed range depends on the current range selected. / The query returns the set current protection level.
Parameters	Name	Type	Range	Default
	<current>	NR2	0.0-10.0	0.1
	<low range>	NR2	0.0-1.0	0.1
Return	Returns the set OCP level as float in Amperes.
Usage example	Set the OCP level to 2.0A: / CURR:PROT 1.5
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]CURRENT:RANGe

### [SOURce:]CURRent:RANGe
Syntax	[SOURce:]CURRent:RANGe {<currentRange>} / [SOURce:]CURRent:RANGe?
Description	This command sets the desired current range. The range can be specified explicitly as low (L) or high (H) or given as a number. In case of a number de range is set to fit the number. / Setting the current range is only allowed when the state is OFF. / The query returns the set current range (L or H).
Parameters	Name	Type	Range	Default
	<currentRange>	Discrete|NR2	LOW|HIGH|0-10.0	HIGH
Return	Returns the set current range.
Usage example	Set the current range to LOW (because 0.5 fits within low range): / CURR:RANGE 0.5
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]CURRENT / [SOURce:]INPut[:STATe]

## POWER subsystem
### [SOURce:]POWer[:LEVel][:IMMediate][:AMPLitude]
Syntax	[SOURce:]POWer[:LEVel][:IMMediate][:AMPLitude] {<power>} / [SOURce:]POWer[:LEVel][:IMMediate][:AMPLitude]?
Description	This command sets the desired power level in constant power (CP) mode in Watts. / The query returns the set current level.
Parameters	Name	Type	Range	Default
	<power>	NR2	0.0-125.0	10.0
Return	Returns the set power level as float in Watts.
Usage example	Set the mode to CC, set the current level to 1.5A and turn the load on: / INP:MODE CP / POW 50 / INP 1
Errors
Related Commands	[SOURce:]INPut:MODE

### [SOURce:]POWer:PROTection[:LEVel]
Syntax	[SOURce:]POWer:PROTection[:LEVel] {<power>} / [SOURce:]POWer:PROTection[:LEVel]?
Description	This command sets the desired over power protection (OPP) level in Watts. The protection will trigger if the set level is exceeded for the over power delay amount of time. / The query returns the set power protection level in Watts.
Parameters	Name	Type	Range	Default
	<power>	NR2	0.0-125.0	20
Return	Returns the set OPP level as float in Watts.
Usage example	Set the OPP level to 50W: / POW:PROT 50
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]POWer:DELay

### [SOURce]:POWer:PROTection:DELay[:TIMe]
Syntax	[SOURce]:POWer:PROTection:DELay[:TIMe] {<delay>} / [SOURce]:POWer:PROTection:DELay[:TIMe]?
Description	This command sets the over power delay time in seconds. This is the time the set power level has to be exceeded without interruption before the over power protection is triggered. / The query returns the set power protection delay in seconds.
Parameters	Name	Type	Range	Default
	<delay>	NR1	1-600	20
Return	NR1
Usage example	Set the mode to CC, a current of 1A and set OPP to 5W with a delay of 10s and turn the load on: / INP:MODE CC / CURR 1 / POW:PROT 5 / POW:PROT:DEL 10 / INP 1
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]POWer:PROTection[:LEVel]

## RESISTANCE subsystem
### [SOURce]:RESistance[:LEVel][:IMMediate][:AMPLitude]
Syntax	[SOURce]:RESistance[:LEVel][:IMMediate][:AMPLitude] {<resistance>} / [SOURce]:RESistance[:LEVel][:IMMediate][:AMPLitude]?
Description	This command sets the desired resistance level in constant resistance (CR) mode in Ohm. / The query returns the set resistance level in Ohm.
Parameters	Name	Type	Range	Default
	<resistance>	NR2	0.1-100000.0	1000.0
Return	Returns the set resistance level as float in Ohms.
Usage example	Set the mode to CR, set the resistance level to 100Ω and turn the load on: / INP:MODE CR / RES 100 / INP 1
Errors
Related Commands	[SOURce:]INPut:MODE

## VOLTAGE subsystem
### [SOURce:]VOLTage[:LEVel][:IMMediate][:AMPLitude]
Syntax	[SOURce:]VOLTage[:LEVel][:IMMediate][:AMPLitude] {<voltage>} / [SOURce:]VOLTage[:LEVel][:IMMediate][:AMPLitude]?
Description	This command sets the desired voltage level in CV mode. The allowed range depends on the voltage range selected. / The query returns the set voltage level.
Parameters	Name	Type	Range	Default
	<voltage>	NR2	0.0-80.0	10.0
	<low range>	NR2	0.0-10.0	3.3
Return	Returns the set voltage level as float in Amperes.
Usage example	Set the mode to CV, set the voltage level to 5V and turn the load on: / INP:MODE CV / VOLT 5 / INP 1
Errors
Related Commands	[SOURce:]INPut:MODE / [SOURce:]VOLTage:RANGe

### [SOURce:]VOLTage:PROTection[:LEVel]
Syntax	[SOURce:]VOLTage:PROTection[:LEVel] {<voltage>} / [SOURce:]VOLTage:PROTection[:LEVel]?
Description	This command sets the desired over voltage protection (OVP) level. The allowed range depends on the voltage range selected. / The query returns the set voltage protection level.
Parameters	Name	Type	Range	Default
	<voltage>	NR2	1.0-85.0	40.0
	<low range>	NR2	1.0-10.5	10.5
Return	Returns the set OVP level as float in Volts.
Usage example	Set the OVP level to 20V: / VOLT:PROT 20
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]VOLTage:RANGe

### [SOURce:]VOLTage:RANGe
Syntax	[SOURce:]VOLTage:RANGe {<voltageRange>} / [SOURce:]VOLTage:RANGe?
Description	This command sets the desired voltage range. The range can be specified explicitly as low (L) or high (H) or given as a number. In case of a number de range is set to fit the number. / Setting the voltage range is only allowed when the state is OFF. / The query returns the set voltage range (L or H).
Parameters	Name	Type	Range	Default
	<voltageRange>	Discrete|NR2	LOW|HIGH|0-80.0	HIGH
Return	Returns the set voltage range.
Usage example	Set the voltage range to LOW (because 5 fits within low range): / VOLT:RANGE 5
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / [SOURce:]VOLTage / [SOURce:]INPut[:STATe]