# CAPACITY subsystem
The SOURce:CAPacity subsystem controls the capacity measurements and limits. Capacity measurements are accumulated measurements over time (e.g. mAh, Wh) when the load is turned on.

### [SOURce]:CAPacity[:STATe]
Syntax	[SOURce]:CAPacity[:STATe] {<bool>} / [SOURce]:CAPacity[:STATe]?
Description	This command enables or disables the capacity measurements and limits when the load is in the ON state. Capacity measurement always stops when the load is in the OFF state. / When this state is OFF the capacity limits are also not enforced.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	ON
Return	The query command returns 0 if the measurement state is enabled, and 1 if the measurement state is enabled, independent of the state of the load.
Usage example	Set capacity measurement status to enabled and check it: / CAP 1 / CAP? / 1
Errors
Related Commands	[SOURce:]INPut

### [SOURce]:CAPacity:LIMit[:ENable]
Syntax	[SOURce]:CAPacity:LIMit[:ENable] {<bool>} / [SOURce]:CAPacity:LIMit[:ENable]?
Description	This command enables or disables the capacity threshold limits checks. When one of the limits is reached the load is turned OFF and the limit protection is set. / The query returns whether the capacity limit check is enabled or disabled. / *RST resets the mode to ON
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	ON
Return
Usage example	CAP? / 1
Errors
Related Commands

### [SOURce]:CAPacity:LIMit:CLEar/TRIPped?
Syntax	[SOURce]:CAPacity:LIMit:CLEar / [SOURce]:CAPacity:LIMit:TRIPped?
Description	This command clears the capacity limit protection if set. / The query returns whether the capacity limit is set (tripped).
Return	The query returns 1 when a capacity limit is reached, 0 otherwise.
Usage example	CAP 1 / CAP:LIM 1 / CAP:LIM? / 0
Errors
Related Commands	[SOURce]:CAPacity:LIMit

### [SOURce]:CAPacity:ZERO
Syntax	[SOURce]:CAPacity:ZERO
Description	This command set all capacity measurements to 0.
Return	None
Usage example	CAP:ZERO
Errors
Related Commands	[SOURce]:INPut:CAPacity 1

### [SOURce]:CAPacity:LIMit:*:STOP
Syntax	[SOURce]:CAPacity:LIMit:AH[:STOP] {<AhStop>} / [SOURce]:CAPacity:LIMit:AH[:STOP]? / [SOURce]:CAPacity:LIMit:TIMe[:STOP] {<TimeStop>} / [SOURce]:CAPacity:LIMit:TIMe[:STOP]? / [SOURce]:CAPacity:LIMit:VOLTage[:STOP] {<VoltStop>} / [SOURce]:CAPacity:LIMit:VOLTage[:STOP]? / [SOURce]:CAPacity:LIMit:WH[:STOP] {<WhStop>} / [SOURce]:CAPacity:LIMit:WH[:STOP]?
Description	These commands set the capacity limits for Ah, Time, Voltage and Wh. For Ah, Time and Wh these limit specify a max value. If any of these values is exceeded the capacity limit state is set. The voltage is specified as a minimum (under voltage). If the voltage drops below the given value the capacity limit it triggered. / The qeuries return the set limit of the corresponding value
Parameters	Name	Type	Range	Default	Unit
	<AhStop>	NR2	0.001-3600.0	10.0	Ah
	<TimeStop>	Integer	1-864000	86400	s
	<VoltStop>	NR2	0.5-80	3.0	V
	<WhStop>	NR2	0.001-3600	10	Wh
Return	The request return the set values for the corresponding limit.
Usage example	Set the capacity limit to max. 100mAh, max. 1 hour, max. 250mWh and min. 3.1V and activate the capacity limit enforcement and the capacity measurements. / CAP:LIM:AH 0.1 / CAP:LIM:TIM 3600 / CAP:LIM:WH 0.25 / CAP:LIM:VOLT 3.1 / CAP:LIM 1 / CAP 1
Related Commands

### [SOURce]:INPut:PROTection:WDOG:CLEar/TRIPped?
Syntax	[SOURce]:INPut:PROTection:WDOG:CLEar / [SOURce]:INPut:PROTection:WDOG:TRIPped?
Description	This command clears the latched protection status that disables the output when a SCPI watchdog timeout is reached. Note the watchdog is NOT turned of and will trigger again if activity or PET commands are not detected within the set delay. / The query returns whether the SCPI watchdog condition is set.
Usage example	The following command clears the SCPI watchdog status: / INP:WDOG:CLE
Related Commands	[SOURce]:INPut:PROTection:WDOG / [SOURce]:INPut:PROTection:TYPe / [SOURce]:INPut:PROTection:DELay

### [SOURce]:INPut:PROTection:WDOG:DELay
Syntax	[SOURce]:INPut:PROTection:WDOG:DELay {<delay>} / [SOURce]:INPut:PROTection:WDOG:DELay?
Description	This command set the SCPI watchdog timeout delay in seconds. / The query returns the current set delay in seconds.
Parameters	Name	Type	Range	Default
	<delay>	Integer	0-3600	10
Return	The query command returns the number of seconds the delay is set to, independent of the status of the watchdog.
Usage example	INP:WDOG:DEL 5 / INP:WDOG:DEL? / 5
Related Commands	INP:WDOG

### [SOURce]:INPut:PROTection:WDOG:PET
Syntax	[SOURce]:INPut:PROTection:WDOG:PET
Description	This command clears the SCPI watchdog inactivity timer.
Return	None
Usage example	INP:PROT:WDOG:PET

### [SOURce]:INPut:PROTection:WDOG:TYPe
Syntax	[SOURce]:INPut:PROTection:WDOG:TYPe {<wdogType>} / [SOURce]:INPut:PROTection:WDOG:TYPe?
Description	This command set the SCPI watchdog type to one of: / Activity: The SCPI watchdog timer is reset on reception of any SCPI command line / PET: The SCPI watchdog timer is only reset on the reception of the INPut:WDOG:PET command.
Parameters	Name	Type	Range	Default
	<wdogType>	Discrete	ACTivity|PET	ACT
Return	The type of SCPI watchdog type. ACT for Activity or PET for PET.
Usage example	Set the SCPI watchdog type to PET, set the delay to 5s, activate the SCPI watchdog: / INP:WDOG:TYP PET / INP:WDOG:DEL 5 / INP:WDOG 1 / Send the PET command at least once every 5 seconds: / INP:WDOG:PET / INP:WDOG:PET / Disable the SCPI watchdog when the critical routine is finished or when the remote program ends: / INP:WDOG 1
Errors
Related Commands	[SOURce:]INPut:WDOG