# SENSE subsystem
The SENSE subsystem controls how measurements and are performed.
### [SENSe:]NPLCycles
Syntax	[SENSe:]NPLCycles {<number>} / [SENSe:]NPLCycles?
Description	This command set the number of power line cycles (50Hz or 60Hz) all measurement will be averaged over.
Parameters	Name	Type	Range	Default
	<number>	NR1	1-100	25
Return	The query command returns the current number of power line cycles used to average measurements.
Usage example	Set power line frequency to 50Hz and the averaging to 10, resulting in an averaging period of 200ms (5 measurement per second): / PLFREQ 50 NPLC 10
Errors
Related Commands	[SENSe:]PLFreq

### [SENSe:]PLFreq
Syntax	[SENSe:]PLFreq {<PLFreq>} / [SENSe:]PLFreq?
Description	This command set the power line frequency used. This is used to match the measurement averaging on this frequency to cancel out power line intereference as much as possible.
Parameters	Name	Type	Range	Default
	<PLFreq>	Discrete	50|60	50
Return
Usage example	Set the power line frequency used for averaging measurments to 50Hz: / PLFREQ 50
Errors
Related Commands	[SENSe:]NPLCycles

### [SENSe]:*:STATistics[:ON]
Syntax	[SENSe]:CURRent:STATistics[:ON] {<bool>} / [SENSe]:CURRent:STATistics[:ON]? / [SENSe]:VOLTage:STATistics[:ON] {<bool>} / [SENSe]:VOLTage:STATistics[:ON]? / [SENSe]:POWer:STATistics[:ON] {<bool>} / [SENSe]:POWer:STATistics[:ON]?
Description	These commands toggle the active tracking of statistics (min, max, avg and count) of different measured values. Statistics can be turned on or off independently for the different measurement. When turned off the values will be held. Turning ON will continue to add to the existing measurements. / The queries returns whether the statistics are ON (1) or OFF (0) for the corresponding measurement.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|1|0	0
Return	The request returns 1 for ON and 0 for OFF.
Usage example	Turn on capacity statistics measurements for current: / CURR:STAT ON
Related Commands	[SENSe:]*:STATistics:CLEAR / FETCh:[SCALAR:]*[:DC]:STATistics?

### [SENSe]:*:STATistics:CLEar
Syntax	[SENSe]:CURRent:STATistics:CLEar / [SENSe]:VOLTage:STATistics:CLEar / [SENSe]:POWer:STATistics:CLEar
Description	This command clears the statistics for the corresponding measurement. All values are set to 0.
Usage example	The following command clears the current statistics: / CURR:STAT:CLE
Related Commands	[SENSe:]*:STATistics[:ON] / FETCh:[SCALAR:]*[:DC]:STATistics?

### [SENSe]:VOLTage:REMote
Syntax	[SENSe]:VOLTage:REMote {<bool>} / [SENSe]:VOLTage:REMote?
Description	This command enables or disables the remote voltage sensing port. Remote sensing can only be changed if the sate of the load is OFF. / The query returns the current status of remote sensing.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|1|0	OFF
Return	The query command returns 1 if remote sensing is ON and 0 if remote sensing is OFF.
Usage example	Turn remote voltage sensing on: / VOLT:REM 1
Related Commands

### [SENSe]:VOLTage:REMote:ERRor[:TRIPped]?
Syntax	[SENSe]:VOLTage:REMote:ERRor[:TRIPped]?
Description	This query returns whether a remote voltage sensing was triggerd when the load was on. This causes a protection status.
Return	1 if a remote voltage sensing error was detected, 0 otherwise
Usage example
Errors	None
Related Commands	[SENSe:]VOLTage:REMote / [SOURce:]INPut:PROTection:CLEar