# SYSTEM subsystem
The SYSTEM subsystem controls generic system settings.
### SYSTem:BEEP
Syntax	SYSTem:BEEP [<duration>]
Description	This command triggers the system beeper for the given duration in ms. If no value is specified the set duration (default 300ms) is used.
Parameters	Name	Type	Range	Default
	<duration>	NR1	0-2500	300
Return	None
Usage example	Beep for 0.5s: / SYST:BEEP 500
Errors
Related Commands	SYSTem:BEEP:DURation

### SYSTem:BEEP:DURation
Syntax	SYSTem:BEEP:DURation <duration> / SYSTem:BEEP:DURation?
Description	This command set beeper duration for all following beeper activation without an explicit duration specified. / The query returns the current set duration.
Parameters	Name	Type	Range	Default
	<duration>	NR1	0-3500	300
Return	Duration in milliseconds.
Usage example	Set the beeper duration to 0.1s for all following beeps and beep: / SYST:BEEP 100 / SYST:BEEP
Errors
Related Commands	SYSTem:BEEP

### SYSTem:BEEP:ON:*
Syntax	SYSTem:BEEP:ON:PROTection {<bool>} / SYSTem:BEEP:ON:PROTection? / SYSTem:BEEP:ON:CAPacity {<bool>} / SYSTem:BEEP:ON:CAPacity? / SYSTem:BEEP:ON:POLarity {<bool>} / SYSTem:BEEP:ON:POLarity? / SYSTem:BEEP:ON:SENSe {<bool>} / SYSTem:BEEP:ON:SENSe? / SYSTem:BEEP:ON:ENCoder {<bool>} / SYSTem:BEEP:ON:ENCoder?
Description	These commands toggle whether the system beeper sounds in certain situations or not. By default all are ON. / The queries return the current status if the beeper sound configuration is ON (1) or OFF (0). / There are different settings for: / PROTection: Beep when protection state is triggered or not. / CAPacity: Beep when the capacity limit state is triggered or not. / POLarity: Beep when a reverse polarity situation occurred or not. / SENSE: Beep when a sense error occurs. / ENCoder: Beep shortly when encoder is rotated or not.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|1|0	1
Return	The request returns 1 for ON and 0 for OFF.
Usage example	Turn off beeper sound on encoder movement: / SYST:BEEP:ON:ENC OFF
Related Commands

### SYSTem:FAN:MODE
Syntax	SYSTem:FAN:MODE {<fanmode>} / SYSTem:FAN:MODE?
Description	This command set the FAN mode. The modes are AUTO and MANUAL. In auto mode the fan speed is automatically set based on the measured temperature of the heatsink according to a lookup table. In manual mode the Fan speed is fixed based to a given percentage. / The query returns the currently set fan mode.
Parameters	Name	Type	Range	Default
	<fanmode>	Discrete	AUTO|MAN	AUTO
Usage example	Set the fan mode to manual and the speed to maximum: / SYST:FAN:MODE MAN / SYST:FAN:SPEED 100
Related Commands	SYSTem:FAN:SPEED / SYSTem:FAN:RPM? / SYSTem:TEMPerature?

### SYSTem:FAN:SPEED
Syntax	SYSTem:FAN:SPEED {<speed>} / SYSTem:FAN:SPEED?
Description	This command set the FAN speed percentage when the fan is in manual mode. / The query returns the currently set fan speed.
Parameters	Name	Type	Range	Default
	<speed>	NR1|Discrete	0 to100,MIN|MAX|DEF	50
Usage example	Set the fan mode to manual and the speed to maximum (100%): / SYST:FAN:MODE MAN / SYST:FAN:SPEED 100
Related Commands	SYSTem:FAN:MODE

### SYSTem:FAN:RPM?
Syntax	SYSTem:FAN:RPM?
Description	This query returns the last measured rpm of the specified fan. If no van is specified the rpm of the heatsink fan is returned. / Note: fan option is not implemented yet. Query always returns the heatsink fan RPM.
Return	The query command returns the last measured rpm value as integer.
Usage example
Related Commands	SYSTem:FAN:MODE / SYSTem:FAN:SPEED

### SYSTem:TEMPerature?
Syntax	SYSTem:TEMPerature?
Description	This query returns the last measured temperature at the heatsink.
Return	Temperature in degree Celsius
Usage example	SYST:TEMP? / 24.3
Errors	None
Related Commands

### SYSTem:TEMPerature:PROTection[:LEVel]
Syntax	SYSTem:TEMPerature:PROTection[:LEVel] {<temp>} / SYSTem:TEMPerature:PROTection[:LEVel]?
Description	This command sets the desired over temperature protection (OTP) level in degree Celsius. The protection will trigger if the set level is exceeded for the over temperature delay amount of time. / The query returns the set temperature protection level in degrees Celcius.
Parameters	Name	Type	Range	Default
	<temp>	NR2|Discrete	10.0 to110.0, MIN|MAX|DEF	80
Return	Returns the set OTP level as float in degree Celcius
Usage example	Set the OTP level to 50°C: / SYST:TEMP:PROT 50
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / SYSTem:TEMPerature:DELay

### SYSTem:TEMPerature:PROTection:DELay[:TIMe]
Syntax	SYSTem:TEMPerature:PROTection:DELay[:TIMe] {<delay>} / SYSTem:TEMPerature:PROTection:DELay[:TIMe]?
Description	This command sets the over temperature delay time in seconds. This is the time the set temperature level has to be exceeded without interruption before the over power protection is triggered. / The query returns the set temperature protection delay in seconds.
Parameters	Name	Type	Range	Default
	<delay>	NR1|Discrete	1 to 600, MIN|MAX|DEF	20
Return	NR1
Usage example	Set the OTP level to 50°C and trigger after 10seconds: / SYST:TEMP:PROT 50 / SYST:TEMP:PROT:DEL 10
Errors
Related Commands	[SOURce:]INPut:PROT:CLEAR / SYSTem:TEMPerature:PROTection[:LEVel]