# CALibration subsystem
The CALibration subsystem handles all calibration related functionality.
## CALibration[:MODE]
Syntax	CALibration[:MODE] {<bool>} / CALibration[:MODE]?
Description	This command enables or disables the calibration mode. All other calibration commands are only valid when in calibration mode. / If calibration mode exits, a *RST is executed.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	The query command returns 0 if the calibration mode is OFF, and 1 if the calibration mode is ON.
Usage example	Enter calibration mode: / CAL 1
Errors
Related Commands

## CALibration:INIT:EEPROM
Syntax	CALibration:INIT:EEPROM
Description	This command initializes the EEPROM with a default set of uncalibrated data. / ALL EXISTING CALIBRATION DATA WILL BE ERASED FROM THE EEPROM / This is needed before individual calibration data set can be written to the EEPROM. Typically this is only done once during initial calibration of the device. / Note: This command takes several seconds to execute and the SCPI interface will not respond during this period.
Parameters	Name	Type	Range	Default
	<mode>	Discrete	CC|CV|CR|CP|DVM|SHORT	CC
Return	None
Usage example	CAL:INIT:EEPROM
Errors
Related Commands	CALibration[:MODE]

## CALibration:TYPe
Syntax	CALibration:TYPe {<caltype>} / CALibration:TYPe?
Description	This command select the range to be calibrated.
Parameters	Name	Type	Range	Default
	<caltype>	Discrete	IMONH|IMONL|UMONH|UMONL|ISETH|ISETL|USETH|USETL|VONH|VONL|OCPH|OCPL|OVPH|OVPL	-
Return	The command doesn’t return anything.
Usage example	Enter calibration mode and switch to calibrating IMONLow range. / CAL 1 / CAL:TYPE IMONL
Errors
Related Commands	CALibration[:MODE] / CALibration:POInt / CALibration:SET

## CALibration:POInt
Syntax	CALibration:POInt {<calpoint>} / CALibration:POInt? / CALibration:POInt:MAX?
Description	This command selects the calibration point within the active calibration range (type). In current implementation there are 2 calpoints per range. / The POInt query returns the currently active calibration point. / The MAX query returns the maximum datapoints in this range. In current implementation this is always 2.
Return	This query returns 0 if a protection is not set, or 1 when it was not detected.
Usage example	Enter calibration mode, switch to calibrating IMONLow range and select first datapoint: / CAL 1 / CAL:TYPE IMONL / CAL:POINT 1
Errors
Related Commands	CALibration[:MODE] / CALibration:POInt / CALibration:SET

## CALibration:SET
Syntax	CALibration:SET {<setpoint>} / CALibration:SET?
Description	This command sets the ‘set’ point within a datapoint. For ranges that read values this is the value that is externally applied as a reference. For ranges that set values this is the raw value that is put into the DAC (rounded to an integer). / The query returns the current set point.
Parameters	Name	Type	Range	Default
	<setpoint>	Float	Depends on range	-
Return	Returns the set value of the datapoint as a float.
Usage example	Enter calibration mode, switch to calibrating UMONLow range, select first datapoint and set the set point to the externally applied voltage of 1.1V: / CAL 1 / CAL:TYPE UMONL / CAL:POINT 1 / CAL:SET 1.1
Related Commands	CALibration[:MODE] / CALibration:POInt

## CALibration:ADC?
Syntax	CALibration:ADC? {<ADClimit>}
Description	Query the ADC min or max values for this calibration point. This is only applicable for ranges with and ADC (current and voltage monitoring).
Parameters	Name	Type	Range	Default
	<ADClimit>	Discrete	MIN|MAX	-
Returns	ADC value as an integer.
Usage example	Enter calibration mode, switch to calibrating UMONLow range, select first datapoint and query the maximum value of the ADC: / CAL 1 / CAL:TYPE UMONL / CAL:POINT 1 / CAL:ADC? MAX / 8388607
Related Commands

## CALibration:MEASure
Syntax	CALibration:MEASure / CALibration:MEASure?
Description	Perform a calibration measurement. / For range that measures values: The ADC value is read. / For ranges that set values directly with a DAC: The DAC value is set. The real value should be read externally and set with CAL:SET. / For ranges that set values to a trigger level: A search is started to find the DAC value corresponding to the externally applied level. The applied level should be set with SET. / After the measure function, the calibration values for this range can be stored in the EEPROM with CAL:SAVE. / The result from the measurement can be obtained with the MEASure? query.
Return	The query command returns the measured value (from the ADC or DAC).
Usage example	Enter calibration mode, switch to calibrating UMONLow range, select first datapoint, set the set point to the externally applied voltage of 1.1V and performa a calibration measurement: / CAL 1 / CAL:TYPE UMONL / CAL:POINT 1 / CAL:SET 1.1 / CAL:MEAS
Related Commands

## CALibration:MEASure:RUNning?
Syntax	CALibration:MEASure:RUNning?
Description	Query returns whether a cal:meas command is currently executing.
Return	Boolean. 1 = calibration measurement is running, 0 = no calibration measurement is running.
Usage example	CALibration:MEASure:RUNning? / 1

## CALibration[:TYPe]:SAVE
Syntax	CALibration[:TYPe]:SAVE
Description	This command saves the currently active datapoints for the active range to the EEPROM.
Return	None
Usage example	Enter calibration mode, switch to calibrating UMONLow range, select first datapoint, set the set point to the externally applied voltage of 1.1V and perform a calibration measurement: / CAL 1 / CAL:TYPE UMONL / CAL:POINT 1 / CAL:SET 1.1 / CAL:MEAS / Switch to second point, set point to externally applied voltage of 9.51V and perform a calibration measurement: / CAL:POINT 2 / CAL:SET 9.5 / CAL:MEAS / Save the set and measured values for point 1 and point 2 of UMONL to the EEPROM: / CAL:SAVE
Errors
Related Commands

## CALibration[:TYPe]:RESET
Syntax	CALibration[:TYPe]:RESET
Description	This command resets the values for this calibration range to the values stored in the EEPROM.
Return	None.
Usage example
Errors
Related Commands